Antibacterial Properties and Cytocompatibility of

Tantalum Oxide Coatings with Different Silver

Content

Running title: Antibacterial Tantalum Oxide Coatings

Running Authors: Heng-Li Huang et al.

Heng-Li Huang

School of Dentistry, China Medical University, Taichung 404, Taiwan

Yin-Yu Chang a)*

Department of Mechanical and Computer-Aided Engineering, National Formosa University,

Yunlin 632, Taiwan

Hung-Jui Chen

Department of Mechanical and Computer-Aided Engineering, National Formosa University,

Yunlin 632, Taiwan

Yu-Kai Chou

Department of Mechanical and Computer-Aided Engineering, National Formosa University,

Yunlin 632, Taiwan

Chih-Ho Lai

School of Medicine, China Medical University, Taichung 404, Taiwan

Michael YC Chen

Division of Oral & Maxillofacial Surgery, China Medical University Hospital, Taichung 404, Taiwan

a) Electronic mail: yinyu@mail2000.com.tw

Tantalum (Ta) oxides and their coatings have been proved to increase their applications

in the biomedical fields by improving osseointegration and wear resistance. In this study,

Ta oxide coatings containing different proportions of Ag are deposited on SS304

materials. A twin-gun magnetron sputtering system is used to deposit the tantalum oxide2

Ag coating. In this study, S. aureus, which exhibits physiological commensalism on the

human skin, nares, and mucosal and oral areas, is chosen as the model for in vitro antibacterial

analyses via a fluorescence staining method employing Syto9. The

cytocompatibility and adhesive morphology of human skin fibroblast cells (CCD-966SK)

on the coatings are also determined by using the MTT assay. This study shows that Ta2O5

and Ta2O5-Ag coatings with 12.5 at.% of Ag exhibit improved antibacterial effects

against S. aureus and have good SKF cell cellular biocompatibility.

I. INTRODUCTION

Materials used in medical devices, particularly in those applications in which the

devices contact or are implanted in the body, are typically described as biomaterials and

possess unique requirements. Tantalum (Ta), despite its costliness, has been suggested in

addition to Ti as a potential alternative for orthopedic and dental applications in recent

years due to its excellent corrosion resistance, fracture toughness, and biocompatibility.1–

3 Previous studies have shown that Ta metal is applicable for osteogenesis in animal

implantation tests, and suitable for cell adhesion, proliferation and differentiation in invitro

studies.3–7 The biocompatibility of Ta has been shown, and it has since been widely

applied to various biomedical materials and devices.

Tantalum composites show different material characteristics depending on the surface

modification technique and the chemical bonding structures.3,8,9 A thin layer of tantalum

oxide, which naturally exists on the surface of Ta metal, may play a key role in its

biocompatibility. Inorganic tantalum oxide is radio-opaque, chemically and thermally

3

stable, and possesses good cytocompatibility, with better corrosion resistance than

untreated Ti surfaces.10

It has been reported that Ag-doped TaN and Cu-doped TaN with nanoparticles can

decrease the multiplication of Escherichia coli. (E. coli.) bacteria, and exhibits improved

antibacterial effect11–13. Silver is known to be an effective antibacterial agent due to its

specific antimicrobial activity.14,15 Previous studies have demonstrated that Ag-doped

TaN and TiO2 can decrease the multiplication of E. coli and S. aureus bacteria, and

exhibit improved antibacterial effects.16–18 The present study investigates the antibacterial

performance of Ta2O5-Ag coatings and its cytocompatibility with human skin fibroblast

cell proliferation activity.

II. EXPERIMENTAL

Tantalum, Ta2O5 and Ta2O5-Ag coatings were deposited on SS304 stainless steel

samples [surface roughness (Ra)=0.2 m] using pulsed unbalanced magnetron sputtering

with high-purity Ta and Ag targets (99.99 at.%). To enhance film adhesion, Ta ion

bombardment (Ta cathode power of 50 W and bias voltage of −600V) was employed

prior to deposition. After etching, the Ta coating was deposited using a Ta cathode power

of 200 W, and with a bias voltage of −80 V. The Ta2O5 and Ta2O5-Ag coatings were

deposited by driving the cathodes in pulsed DC mode at a frequency of 20 kHz. The

Ta2O5 was deposited with the cathode power of the Ta target set to 100 W at 0.5 Pa.

In order to prepare Ta2O5-Ag with different Ag contents, the cathode power of Ta

was kept at 100 W while the cathode power of Ag was set at the two values of 20 and 40

W. A substrate bias voltage of −80 V was used. The total thickness of the coatings was

4

controlled to be 0.6~0.7 μm by using a deposition time of 20 min. The temperature of the

sample during the deposition, as measured by a thermocouple, was kept within the range

of 110±20C. After deposition, the deposited Ta2O5-Ag coatings were taken from the

coating chamber and annealed at 700°C for 30 seconds using a rapid thermal annealing

(RTA) system.

The surface morphology of the deposited coatings was examined using highresolution

field-emission scanning electron microscopy (FESEM, Joel JSM-7000F, Jeol

Co., Mingdao University, Taiwan). The composition of the deposited films was

determined by an energy dispersive spectroscope (EDS, Joel JSM-7000F, Jeol Co.,

Mingdao University, Taiwan) and their chemical-binding characteristics were found via

X-ray photoelectron spectroscopy (XPS; PHI1600, Perkin Elmer Inc., National Tsing

Hua University, Taiwan) with Mg Kα radiation. Survey spectra in the range of 0–1000

eV were recorded for each sample, followed by high-resolution spectra of different

elemental peaks, from which the composition was calculated. The spectral ranges of

26±10, 530±15 and 370±15 eV corresponded to the binding energies of Ta 3d, O 1s and

Ag 3d, respectively. The energy was calibrated with reference to the Au 4f7/2 peak from a

clean gold surface at 83.8 eV. Glancing-angle X-ray diffractometry (XRD, PANalytical

X’pert Pro, PANalytical Co., Mingdao University, Taiwan) was employed at a glancing

angle of 1° for phase identification. The static contact angle was measured, and the

obtained images were analyzed to calculate the contact angle of deionized water for each

sample at room temperature (24~26°C).

The retention of bacteria on the coated samples was determined by a fluorescence

staining method employing SYTO9 nucleic acid stain. First, 500 μl of S. aureus

5

suspension (1×105 cfu/ml) was added to the sample surface. After incubation for 6 h at

37°C under a relative humidity of 96% and avoiding light exposure, the sample surfaces

were rinsed three times with phosphate-buffered saline, and then the retained bacteria

were fixed with 4% paraformaldehyde (Sigma-Aldrich, St. Louis, MO, USA) and stained

with 10 M STYO9 for 30 min at room temperature (24~26°C). The bacteria that were

adhered to the samples were quantified by measuring the fluorescence detected at 488 nm

by an enzyme-linked immunosorbent assay reader. The results were determined in three

independent experiments performed in duplicate, and quantified in units of relative

fluorescence intensity.

The proliferation of a human skin fibroblast (SKF) cell line (CCD-966SK) was

examined with an MTT assay (Sigma-Aldrich) after the cells were cultured on an

uncoated SS304 stainless steel surface, and SS304 coated with Ta2O5 and Ta2O5-Ag. The

substance used for the MTT assay was a 3-(4,5-dimethylthiazol-2-yl)-2,5-

diphenyltetrazolium (MTT) salt, which turns into a purple formazan product in the

presence of viable mitochondria in living cells. The SKF cells (3 ml) were seeded and

incubated at 37°C in 5% carbon dioxide (CO2) for 48 h, at which time proliferation was

achieved. The MTT (5 mg/ml) was added to the cultured cells and incubated for a further

2 h. The formazan was eluted using 100 μl of isopropanol. The absorbance of the purple

formazan was quantified as the optical density (OD) measured at 570 nm by a

SpectraMax spectrophotometer with SoftMax Pro 5.2 241 software. The OD of formazan

reflects the level of cell metabolic activity, with higher OD values indicating a larger

number of living cells on the sample, and hence better biocompatibility.

6

Experiments were repeated independently in duplicate. The statistical correlation

of the results of antibacterial activity tests between the coated samples and uncoated

SS304 plates was determined via Student’s t-test. Differences were considered significant

at the p<0.05 level.

III. RESULTS AND DISCUSSION

A. Microstructure characterization

The elemental composition of the deposited Ta2O5 coating is comprised of 31.3

at.% Ta and 68.7 at.% O. A Ta2O5-Ag coating is formed by cosputtering Ta and Ag with

O2 introduced during the coating process. The use of different Ag cathode powers (20

and 40 W) produces Ta2O5-Ag coatings containing 12.5 at.% and 23.3 at.% Ag, which

are denoted as Ta2O5-Ag20W and Ta2O5-Ag40W, respectively.

Previous studies measured the binding energy of Ta 4f7/2 for the deposited Ta2O5

coatings to be 26.5–27 eV.19 The theoretical binding energy for O 1s in Ta2O5 is typically

530–532 eV, and similar values are observed for our deposited Ta2O5 coatings. The

bonding states of the deposited Ta2O5-Ag40W coatings are also characterized by XPS.

Figure 1 shows the Ta 4f, O 1s and Ag 3d XPS spectra of the deposited Ta2O5-Ag40W

coating. The Ta 4f spectrum indicates a major Ta 4f7/2 peak centered at 26.9 eV, which is

attributable to Ta2O5, while the Ta 4f7/2 peak shoulder centered at 23.1 eV is attributable

to TaO.19,20 The Ag 3d5/2 spectrum shows a peak at 367.7 eV, which is attributable to

Ag2O. The binding energy for O 1s in the deposited Ta2O5-Ag40W is 531.3 eV, which

corresponds to typical Ta oxides and Ag2O. The results suggest that the deposited Ta2O5-

Ag40W coatings are mainly comprised of Ta2O5 and silver oxides.

7

Figure 2 shows SEM micrographs of the deposited Ta, Ta2O5, Ta2O5-Ag20W and

Ta2O5-Ag40W coatings. A crystalline morphology is observed for Ta, while the Ta2O5

exhibits a denser structure. A recent study by Hsieh et al.21 showed that Ag particles were

formed on nitride coatings by using RTA to improve the antibacterial performances.

After RTA, Ag particles would agglomerate inside the coating matrix and on the film

surface. In our present study, the deposited Ta2O5-Ag coatings are annealed at 700°C for

30 seconds by RTA. As shown in Fig. 2, many particles with diameters of several tens of

nm and up to 160 nm are distributed in the annealed Ta2O5-Ag20W and Ta2O5-Ag40W

coatings.

Looking at the XRD measurements, Ta2O5, Ta2O5-Ag20W and Ta2O5-Ag40W

coatings display a broad diffraction crest that appears in the pattern at approximately 2θ

of 20–40°, and they all exhibit an amorphous nature. A previous study by Y.S. Sun et

al.22 showed that amorphous Ta2O5 exhibited good resistance to corrosion in simulated

body plasma (SBP) and improved the surface wettability and cell adhesion of Ti. In the

present study, the inclusion of Ag will influence the surface hydrophobicity of the Ta2O5

coated samples. Figure 3 shows that the contact angle is 86.1±1.7° for uncoated SS304

(control) and 73.1±6.1° when coated with Ta. The SS304 coated with Ta2O5 and Ta2O5-

Ag20W exhibit contact angles similar to the uncoated SS304 (85.8±1.2° and 88.6±1.6°,

respectively), while the Ta2O5-Ag40W coating with 23.3 at.% Ag exhibits the highest

contact angle of 94.8±2.7°. These observations indicate that the presence of Ag in the

Ta2O5-Ag coatings results in a higher contact angle than for amorphous Ta2O5 and the

control samples, and it exhibits hydrophobicity.

8

B. Antibacterial performance

The extremely complicated mechanism of bacterial adhesion to surfaces is

affected by many factors, including the environment, bacterial properties, and material

surface characteristics such as the chemical composition, surface charge, hydrophobicity,

and topography. It is particularly important to fabricate a coating material with both

antimicrobial capacity and cytocompatibility to increase the likelihood of implant success.

S. aureus are considered to be common causes of implant-associated infections and periimplantitis.

23,24 Once S. aureus adhere to a metal surface they form a biofilm that is

difficult to treat clinically, since biofilms protect the bacterial community from

competing microorganisms. In the present study, S. aureus are observed on the surfaces

of SS304 (control) and the coated samples after incubation. The measured relative

SYTO9 fluorescence intensities of S. aureus retained on uncoated and the SS304 plates

coated with Ta, Ta2O5 and Ta2O5-Ag are shown in Fig. 4. The fluorescence intensity of

bacterial retention is highest for the uncoated SS304 surfaces. The Ta2O5-coated samples

show lower bacterial retention than the uncoated SS304 surfaces, and incorporating Ag

with the Ta2O5 significantly inhibits S. aureus. There is a significant difference (P<0.01)

between the fluorescence intensity of the Ta2O5 (81.5±5.2) and Ta2O5-Ag20W (81.5±1.6)

samples and that of the Ta2O5-Ag40W sample (73.0±1.1). The fluorescence intensity is

lowest for the Ta2O5-Ag40W surface corresponding to the fewest adhering bacteria,

which is further substantiated by it possessing the highest contact angle and

hydrophobicity, and it shows significant short-term antibacterial effects against S. aureus.

Previous studies have confirmed the antibacterial activities of Ti-Ag, TaN-Ag, TiO2-Ag

and ZrO2-Ag coatings against S. aureus in vitro.13,18,25,26 The antibacterial action of Ag

9

may be attributed to three actions of the Ag ions. First, the ions may uncouple respiratory

electrons transported from oxidative phosphorylation and inhibit the respiratory chain

enzymes. Second, the ions may interfere with the membrane permeability to protons and

phosphate.27 Finally, ions may raise the silver ion-mediated reactive oxygen species

generation to inhibit the bacterial ability.28 In this way, samples with a large content of

Ag particles may release more Ag ions, which can markedly damage existing bacteria.

H.J. Park et. al.28 showed that the antibacterial activity of the superoxide-radicalgeneration

mechanism was closely related to the thiol-interaction mechanism. The

inhibitory effect of Ag ions that are released from Ag under complex physiological

conditions is believed to inactivate bacteria and prevent bacterial replication.29,30 The

deposited Ta2O5-Ag coatings may release silver ions in aqueous solution, which further

enhances the bactericidal activity,31 because Ag ions strongly bind to electron donor

groups in biological molecules containing sulfur, oxygen or nitrogen.32 The antibacterial

properties of surface coatings containing Ag can suppress microbial proliferation and

thereby reduce bacterial counts. This might also decrease the probability of implantrelated

infections.

C. Cytocompatibility and proliferation of SKF cells

The surface texture of the implant materials can be modified to enhance the

tissue-implant attachment, thereby producing a material that is at least as good as SS304.

The present study used an MTT assay to evaluate the viability of SKF cells for the

application of medical devices. We measure the optical density of the formazan produced

by SKF cells (Fig. 5), and find the uncoated SS304 (control) to have an optical density of

10

0.205±0.025. In this study, SS304 plates coated with Ta show a better biocompatibility

than the control, and Ta2O5 shows good viability and proliferation of SKF cells. The

Ta2O5-Ag20W does not differ significantly from uncoated SS304 plates (p>0.05) in

terms of viability and proliferation of SKF cells. Therefore, the MTT assay reveals that

the cytocompatibility is not degraded by modifying the Ta2O5-Ag coated surface with Ag

less than 12.5 at.%. As shown in the SEM micrographs in Fig. 6, adhered SKF cells

exhibit long filopodia that are well attached to the surface of the Ta, Ta2O5 and Ta2O5-

Ag20W coatings. The SEM results reveal that the modified surfaces also have a

comparable biocompatibility because they not only favor cell attachment, but also

maintain a normal spreading architecture of the SKF cells. J. Liao et. al.26 showed that

titanium plates with depositions of nanosilver exhibited good antibacterial activity to S.

aureus and E. coli., and did not show any cytotoxicity to the human gingival fibroblasts.

However, in the present study, the Ta2O5-Ag coating with 23.3 at.% Ag has a lower

optical density (p<0.01), corresponding to a reduced viability and proliferation of SKF

cells compared to the uncoated sample. A previous study by Y. K. Chen et. al.33 showed

that surfaces coated with silver-containing hydroxyapatite possess hydrophilicity that

could be beneficial to the proliferation of cells on its surface. Compared with the Ta2O5-

Ag20W, however, the Ta2O5-Ag40W coating has a higher contact angle (94.8±2.7°) and

exhibits hydrophobicity. The hydrophobic feather of the Ta2O5-Ag40W coatings,

therefore, causes decreased viability and proliferation of SKF cells compared to the

Ta2O5-Ag20W sample. In addition, the higher content of Ag (23.3 at.%) in the Ta2O5-

Ag40W coatings may act as a foreign body, restricting the normal activity of the SKF

cells. It is conceivable that, under such circumstances, the SKF cells are not expected to

11

survive.34 Therefore, the present study reveals the harmful effects of Ta2O5-Ag coatings

with high silver contents for SKF cells, and demonstrates that applications of Ta2O5-Ag

should be given special attention. Both the MTT assay and SEM results suggest that the

Ta2O5-Ag coatings with Ag less than 12.5 at.% do not show detectable cytotoxicity to

SKF cells, and demonstrates a favorable cell adhesion to the coated samples which is

indicative of a progression of the differentiation process.

IV. SUMMARY AND CONCLUSIONS

In this study, a series of Ta, Ta2O5 and Ta2O5-Ag coatings are synthesized using

twin-gun reactive magnetron sputtering. From the XRD measurements, Ta2O5, Ta2O5-

Ag20W and Ta2O5-Ag40W coatings show a broad diffraction crest to appears in the

pattern at approximately 2θ of 20~40°, and exhibit an amorphous structure. The XPS

results suggest that the deposited Ta2O5-Ag40W coatings are mainly comprised of Ta2O5

and silver oxides. The Ta2O5-Ag40W coating with 23.3 at.% Ag exhibits the highest

contact angle of 94.8±2.7°, and displays hydrophobicity. The presence of Ag in the

Ta2O5-Ag coatings results in a higher contact angle than with Ta2O5 alone.

The antibacterial performance of the Ta2O5 and Ta2O5-Ag coatings against grampositive

bacteria (S. aureus), and their cytocompatibility with SKF cells is investigated.

The antibacterial performances are superior in the Ta2O5 and Ta2O5-Ag coatings than the

uncoated and Ta-coated samples. The Ta2O5-Ag coating with 23.3 at.% Ag has a lower

optical density (p<0.01), exhibiting reduced viability and proliferation of SKF cells than

the uncoated sample. The results indicate that Ta2O5 and Ta2O5-Ag coatings with 12.5

12

at.% of Ag exhibit improved antibacterial effects against S. aureus and have satisfactory

SKF cell viability.

ACKNOWLEDGMENTS

The Authors are grateful for funding in part from the National Science Council of Taiwan

under the contracts NSC 101-2221-E-150-032 and NSC 101-2622-E-150-001. In

addition, instrument assistance from the Common Lab for Micro/Nano Science and

Technology at the National Formosa University was sincerely appreciated.

1 A. Fukunaga, S.Y. Chu and M.E. McHenry, J. Mater. Res. 13/9, 2465 (1998).

2 N. Wang, H.Y. Li, J.S. Wang, S. Chen, Y.P. Ma and Z.T. Zhang, ACS Appl. Mater.

Interf. 4, 4516 (2012).

3V.K. Balla, S. Bose, N.M. Davies, A. Bandyopadhyay and J. .Miner. Met. Mater. Soc.

62, 61(2010).

4M. Kiuru, E. Alakoski, V.M. Tiainen, R. Lappalainen, A. Anttila and J. Biomed. Mater.

Res. B: Appl. Biomater. 66, 425 (2003).

5C. Balagna, M. Faga and S. Spriano, Mater. Sci. Eng. C 32, 215 (2012).

6K.J. Welldon, G.J. Atkins, D.W. Howie and D.M. Findlay, J. Biomed. Mater. Res. 84,

691(2008).

7V. Sollazzo, F. Pezzetti, L. Massari, A. Palmieri, G. Brunelli, I. Zollino, A. Lucchese, G.

Caruso and F. Carinci, Int. J. Periodontics Restor. Dent. 31, E17 (2011).

8V.K. Balla, S. Banerjee, S. Bose and A. Bandyopadhyay, Acta Biomater. 6 (2010) 2329.

13

9Y.X. Leng, J.Y. Chen, P. Yang, H. Sun, J. Wang and N. Huang, Nucl. Instrum. Methods

Phys. Res. Sect. B-Beam Interact. Mater. Atoms 242, 30 (2006).

10V.A. Shvets, V.Sh. Aliev, D.V. Gritsenko, S.S. Shaimeev, E.V. Fedosenko, S.V.

Rykhlitski, V.V. Atuchin, V.A. Gritsenko, V.M. Tapilin and H. Wong, J. Non-Cryst.

Solids 354, 3025 (2008).

11J.H. Hsieh, C.C. Tseng, Y.K. Chang, S.Y. Chang and W. Wu, Surf. Coat. Technol. 202,

5586 (2008).

12P.C. Liu, J.H. Hsieh, C. Li, Y.K. Chang and C.C. Yang, Thin Solid Films 517, 4956

(2009).

13H.L. Huang, Y.Y. Chang, M.C. Lai, C.R. Lin, C.H. Lai and T.M. Shieh, Surface &

Coatings Technology, 205, 1636 (2010).

14H. Gerspach, S. Stübinger and J. Meyer, Clin. Oral Impl. Res. 21, 277 (2010).

15A. Regiel, S. Irusta, A. Kyzioł, M.Arruebo and J. Santamaria , Nanotechnology 24,

015101 (2013).

16J.H. Hsieh, C.C. Tseng, Y.K. Chang, S.Y. Chang and W. Wu, Surf. Coat. Technol. 202,

5586 (2008).

17Y.Y. Chang, C.H. Lai, J.T. Hsu, C.H. Tang, W.C. Liao and H.L. Huang, Clin. Oral

Invest. 16, 95 (2012).

18H.L. Huang, Y.Y. Chang, J.C. Weng, Y.C. Chen, C.H. Lai and T.M. Shieh, Thin Solid

Films 528, 151 (2013).

19O. Kerrec, D. Devilliers, H. Groult and P. Marcus, Mater. Sci. Eng. B55, 134 (1998).

20M. Khanuja, H. Sharma, B.R. Mehta and S.M. Shivaprasad, J. Electron. Spectrosc.

Relat. Phenom. 169, 41 (2009).

14

21J.H. Hsieh, T.H. Yeh, C. Li, C.H. Chiu and C.T. Huang, Vacuum 87, 160 (2013).

22Y.S. Sun, J.H. Chang and H.H. Huang, Thin Solid Films 528, 130 (2013).

23M. Kronström, B. Svenson, M. Hellman and G.R. Persson, Int. J. Oral Max Impl. 16,

201 (2001).

24E. Barth, Q.M. Myrvik, W. Wagner and A.G. Gristina, Biomaterials. 10, 325(1989).

25B. S. Necula, L. E. Fratila-Apachitei, S.A.J. Zaat, I. Apachitei and J. Duszczyk, Acta

Biomater. 5, 3573 (2009).

26J. Liao, M. Anchun, Z. Zhu and Y. Quan, Int. J. Nanomed. 5, 337 (2010).

27C.N. Lok, C.M. Ho, R. Chen, O.Y. He, W.Y. Yu, H. Sun, P.K.H. Tam, J.F. Chiu and

C.M. Che, J. Proteome. Res. 5,916 (2006).

28H.J. Park, J. Y. Kim, J. Kim, J.H. Lee, J.S. Hahn, M. B. Gu and J. Yoon, Water Res. 43,

1027 (2009).

29G. Colon, B.C. Ward, T.J. Webster. J. Biomed. Mater. Res. A. 595/604, 78 (2006).

30Y. Osamu, K. Miyako, S. Jun, N. Zenbee, J. Mater. Sci. Mater. Med. 847/851, 15

(2004).

31J.R. Morones, J.L. Elechiguerra, A. Camacho and J.T. Ramirez, Nanotechnology 16,

2346 (2005).

32H.Y.Song, K.K. Ko, L.H. Oh, B.T. Lee, Eur. Cells Mater. 11, 58 (2006).

33Y. K. Chen, X. B. Zheng, B. Li, Y. T. Xie,C. X. Ding and S. W. Lee, Mater. Sci. Forum

620/622, 567 (2009).

34P.V. Asharani1, Y. L. Wu, Z. Gong and S. Valiyaveettil, Nanotechnology 19, 255102

(2008).

15

Figure Captions

Figure 1. (Color online) The XPS spectra of the deposited Ta2O5-Ag40W coatings

showing the (a) Ta 4f, (b) O 1s and (c) Ag 3d peaks.

Figure 2. SEM micrographs showing the surface morphologies of the deposited (a) Ta, (b)

Ta2O5, (c) Ta2O5-Ag20W and (d) Ta2O5-Ag40W coatings. The arrows indicate

nanoparticles on the annealed Ta2O5-Ag20W and Ta2O5-Ag40W coatings.

Figure 3. (Color online) Contact angles of the SS304 plates uncoated (control) and coated

with Ta2O5 and Ta2O5-Ag.

Figure 4. (Color online) Relative fluorescence intensities of adherent S. aureus colonies

on the surfaces of uncoated SS304 (control), and SS304 coated with Ta, and

Ta2O5 and Ta2O5-Ag. *Probability values of p<0.05 are considered indicative of

significant differences.

Figure 5. (Color online) Cell viability evaluation using an MTT assay of SKF cells

incubated on the uncoated SS304 surfaces (control) and SS304 surfaces coated

with Ta, Ta2O5 and Ta2O5-Ag. Data shown are the mean and standard deviation

values. *Probability values of p<0.05 were considered indicative of significant

differences.

Figure 6. SEM micrographs of SKF cells incubated on SS304 surfaces coated with (a) Ta,

(b) Ta2O5 and (c) Ta2O5-Ag20W.