Biomechanical investigations of the expanded platform-switching concept in immediately loaded small-diameter implants

Aaron Yu-Jen Wu1, Hsuan Lung1, Heng-Li Huang2, Dr. Chee3*

Short title: Biomechanical analysis of platform-switching small-diameter implants

1 Department of Dentistry, Chang Gung Memorial Hospital & College of Medicine,     Chang Gung University, 123, Ta-Pei Road, Niao-Sung, Kaohsiung 83305, Taiwan.

2 School of Dentistry, China Medical University, 91 Hsueh-Shih Road, Taichung 40402, Taiwan.

3 Ostrow School of Dentistry, University of Southern California, Los Angeles, CA, USA

*Corresponding Author: Ostrow School of Dentistry, University of Southern California, Los Angeles, CA, USA. e-mail: wchee@usc.edu

Abstract

Statement of problem: Using a small-diameter implant may increase the stress of bone around the implant neck, but an expanded platform design may relief them. However, no study has investigated the biomechanical effect of these two kinds of designs within an implant. Purpose: This study evaluated the biomechanical effects of an expanded platform-switching design for immediately loaded small-diameter implants on bone strains. Materials and methods: Three groups of artificial jawbone models were prepared for small-diameter (3.25 mm) and standard-diameter (4.0 mm) implants with expanded or regular platform designs. The platform-switching implant design was implemented by assembling implants with a smaller connected abutment. Specimens were tested under both vertical and lateral static loads at 190 N. Peak values of principal microstrain of bone were recorded and analyzed statistically by Kruskal-Wallis test and multiple-comparisons Bonferroni test. The initial stability of each implant was also measured for three types of implant. Results: Under vertical loading, the bone strain was lowest for the regular type of immediately loaded small-diameter implant. Under lateral loading, the peak bone strain around the platform-switching small-diameter implant was up to 74.9% lower than that of the regular type of small-diameter implant. Increasing the implant diameter from 3.25 mm to 4.0 mm reduced the bone strain by about 10% and 30% under lateral and vertical loading, respectively. The initial implant stability did not varying significantly among the implants. Conclusions: Using the expanded platform-switching design could decrease the marginal bone strains around immediately loaded small-diameter implants only under lateral loading.

Keywords: Small-diameter implant, expanded platform-switching, implant diameter, bone strain

Introduction

Progressive crestal bone loss is one of the leading signs for predicting implant failure after achieving primary stability and osseointegration. How to preserve crestal bone around implants is currently a primary area of research focus. Although the etiological factors underlying bone loss have not been fully established, previous studies have suggested that the main causal factors of marginal bone loss are peri-implantitis and occlusal overload. 1, 2 The biomechanical factors that influence the stress in the bone around an implant include the implant design, diameter, material properties, and surface criteria. 3, 4 Baggi et al. 5 demonstrated that the stress was concentrated in the cortical bone around the implant neck and was influenced primarily by the implant diameter, irrespective of the implant length. Clinically, when the width of the alveolar bone is insufficient or where a concavity of the ridge is present (especially in the anterior area), extensive bone augmentation procedures may be necessary before implant placement. In these circumstances, if the patient does not accept or is not suited for additional surgeries, using a small-diameter implant may be a useful alternative method.

Lazzara and Porter (2006) 6 found some cases in which the long-term preservation of marginal bone was better when using an abutment smaller than the implant neck; this approach came to be known as platform switching. Therefore, peri-implant marginal bone preservation can be more noticeable when the implant and suprastructure have different diameters. 7-9 Platform switching makes it possible to use abutments with a diameter smaller than the implant neck or body width, or alternatively an implant design can be used in which the implant platform is expanded, whereby the implant neck diameter is larger than the diameter of the implant body. 10 In 2009, Hsu et al. used three-dimensional finite-element analysis  to investigate the behaviors of reduced-platform restorations. They found that in all such restorations, the prosthetic loading forces transmitted to the bone–implant interface in the immediately loaded implant model were reduced by 10%. 11

Even though small-diameter implants are useful when the proximity of roots does not allow the safe placement of larger diameter implants without involving adjunctive procedures to move the roots, 12 using a small-diameter implant may increase the  stress and/or strain of bone around the implant neck. The use of an expanded platform design with small-diameter implants may be an option for increasing the implant diameter at the marginal bone area. Unfortunately, there have been few studies of the immediate loading of small-diameter implants, especially those with an expanded platform. In 2010, Cocchetto et al. 8 used both clinical and radiographic evaluations to determine the biologic effects of using an expanded platform-switching restorative protocol in humans. Their results indicated that the crestal bone loss may be lower in appropriately selected patients who receive expanded platform-switching implants than when using non-platform-switching approaches.

Using a smaller diameter implants may increase the bone stresses and/or strains around the implant, which makes it worthwhile to consider modifications to its component design. The use of an expanded platform design with small-diameter implants may be an option for utilizing the platform-switching concept and increasing the implant diameter in the marginal bone area. However, few studies have investigated the biomechanical features of these two kinds of implant designs, and so the aim of the present study was to determine the biomechanical effects of applying an expanded platform-switching design with immediately loaded small-diameter implants; that is, those with diameters smaller than 3.4 mm and that accept interchangeable abutments. The hypothesis tested in this study was that expanding the platform at the marginal bone area will reduce the bone strain.

#### Material and methods

#### Implant design parameters and bone specimen preparation

Two types of 11.5-mm-long Nanotite Certain Prevail implants, with diameters of 3.25 mm and 4.0 mm (IIOS343 and IIOS454, respectively; Biomet 3i, Palm Beach, FL, USA) and an internal-hex connection and expanded platform body were selected for analysis and comparison with a control implant (diameter 3.25 mm, length 11.5 mm; NIOSM311, NanoTite Certain MicroMiniplant, Biomet 3i) with an internal-hex connection and a standard threaded body (Fig. 1). For the IIOS343 and NIOSM311 implants, an 8-mm-long internal hex with a diameter of 3.4 mm, 3.8 mm in platform, and height of 2.0 mm (Certain GingiHue Post, Biomet 3i) was used as an abutment placed on the platform of the implant. For the IIOS454 implant, the same type of internal-hex abutment with a diameter of 4.1 mm, 5.0 mm in platform, and height of 2.0 mm were applied (Fig. 2).

#### A Sawbones model of trabecular bone with a density of 0.4 g/cm3 and elastic modulus of 759 MPa (model 1522-05, Pacific Research Laboratories, Vashon Island, WA, USA) was prepared to attach to a 3-mm-thick commercially available synthetic cortical shell (model 3401-02, Pacific Research Laboratories) with an elastic modulus of 16.7 GPa. The density of trabecular bone used in this study was simulated as type 2 (D2) bone 13 based on the bone-density classification of Misch et al. 14. The thickness of the cortical bone was referred to the study of Hahn 15, which indicated that D2 bone was associated with 2.5–4 mm of cortical bone height. The synthetic bone had a rectangular shape with dimensions of 41 cm × 30 cm × 43.5 cm. Three specimens of artificial foam bone were prepared for each implant system.

Implant stability measurement

The peak Insertion Torque Value (ITV)  was measured using a digital torque meter (TQ-8800, Lulton Electronic Enterprise, Taipei, Taiwan) with the implant inserted into the bone block specimen. In order to simulate the interface condition of an immediately loaded implant, the interface between implant and bone was prepared for contact only. After placing the implant, a wireless resonance frequency analyzer (Osstell ISQ, Osstell AB, Gothenborg, Sweden) was used to measure the implant stability quotient (ISQ). Before the measurement, SmartPegs for theinternal-hex connection of 3i implants (Type 1 and Type 15, Osstell AB) were placed onto the top of the implants. Each peg contains a magnetic material in its upper part. When the probe of the Osstell ISQ instrument is near to the SmartPeg, the peg will be vibrated by magnetic pulses and then Osstell ISQ instrument can detect the resonance frequency and convert it into the ISQ value. For each specimen, four ISQ values were obtained in four directions (i.e., buccal, lingual, mesial, and distal directions) of the implant. After connecting titanium abutments, the implant mobility was measured using the PeriotestTM device (Siemens, Bensheim, Germany). The tip of the measurement device was positioned perpendicularly at 2 mm from the abutment, and it impacted the implant four times per second for 4 seconds 16. Periotest Values (PTV) values were also measured four times for the same four directions in each model.

Strain gauge measurement

A self-developed jig was designed with an adjustable rotational screwing device so that both a vertical load and a 30-degree lingual lateral force could be applied in the experiments. Each loading procedure involved applying a force of 190 N 17 to the cylindrical abutment using a universal testing machine (JSV-H1000, Japan Instrumentation System, Nara, Japan) with a head speed of 1 mm/min (Fig. 3). Rectangular rosette strain gauges (KFG-1-120-D17-11L3M3S, Kyowa, Tokyo, Japan) were attached to the buccal and lingual sides of the crestal cortical region around the implant using cyanoacrylate cement (CC-33A, Kyowa). Signals corresponding to the three independent microstrains εa, εb, and εc measured by the rosette strain gauge were sent to a data acquisition system (NI CompackDAQ, National Instruments, Austin, TX, USA) and analyzed by the associated software (LabVIEW SignalExpress 3.0, National Instruments). After each measurement was repeated three times for each specimen, the maximum (εmax) and minimum (εmin) principal microstrains were obtained as follows:

εmax = 1/2(εa+εc)+1/2√[(εa–εc)2+(2εb–εa–εc)2]                       (1)

εmin = 1/2(εa+εc)−1/2√[(εa–εc)2+(2εb–εa–εc)2]                       (2)

Correlation and statistical analysis

The measured primary implant stability and the peak values of principal microstrains of bone under vertical and lateral loadings for the designed scenarios of three implant systems were summarized as medians and interquartile ranges. The Kruskal-Wallis test and multiple comparisons with the Bonferroni test were used to assess differences. All analyses were performed using a commercial statistical software package (SAS version 9.1, SAS Institute, Cary, NC, USA) with an alpha value of 0.05.

Results

Primary implant stability

The ITV and ISQ values did not differ significantly among the NIOSM311, IIOS343, and IIOS454 implants. However, PTV varied significantly (p<0.05) between the implants, being lowest for IIOS343 (Table 1).

Strain gauge analysis

Under vertical loading, the peak values of principal microstrains of bone (simply referred to as the peak bone strains) around the immediately loaded small-diameter implants differed significantly among the NIOSM311, IIOS343, and IIOS454 implants in both the Kruskal-Wallis test and the multiple-comparisons Bonferroni test (Table 2, Fig. 4). The bone strain was highest for the IIOS343 implant, being at least 31.2% higher on the buccal side and 29.8% higher on the lingual side compared to the other implants.

Under lateral loading, the peak bone strains around the implants differed significantly among the NIOSM311, IIOS343, and IIOS454 implants in the Kruskal-Wallis test (p<0.05) (Table 2, Fig. 5). The bone strains were 9.8% lower on the buccal side and 74.9% lower on the lingual side for the IIOS343 implant than for the NIOSM311 implant. The bone strain was lowest for the IIOS454 implant, being 10.7% lower on the buccal side than for the IIOS343 implant (Table 2, Fig. 5).

Discussion

A systematic review of platform-switching implants found that the survival rate did not differ between platform-switching and conventional implants; however, there was less marginal bone loss over time for the former implants. 18 The reason for this phenomenon is not fully understood, but various biologic and mechanical theories have been proposed. The design of platform switching might increase the distance between the inflammatory-cell infiltrate in the microgap and the crestal bone, thereby minimizing the effect of inflammation on marginal bone remodeling. 6, 19 In addition, based on the results of finite-element analyses, some researchers have suggested that this design reduces the stress at the bone–implant interface and in the crestal region of cortical bone by shifting the stress to cancellous bone during loading. 20-22 If the biologic and mechanical advantages of platform switching really exist, adding the platform-switching design to small-diameter implants should improve the outcomes in clinical applications. However, few studies have investigated small-diameter implants with a platform-switching design. Although one study has investigated implants incorporating the expanded platform-switching concept with a provisional restoration immediately after implant placement, and found limited implant crestal bone loss of 0.97±0.39 mm (mean±SD) after a 5-year follow-up period, 23 that study did not include any small-diameter implants.

In the present study, the bone strain value differed significantly among the three groups of implants under lateral loading. The IIOS343 and NIOSM311 implants both have a small diameter (3.25 mm), but the surrounding bone strain decreased more for the IIOS343 implant that for the NIOSM311 implant, which may be due to the expanded platform-switching design of the former. Because platform switching changes the traditional design of the abutment–implant connection, the stress/strain distributions from the abutment to the implant and from the implant to the bone might be influenced when occlusal loading occurs. The present results indicate that using a platform-switching design for an immediately loaded small-diameter implant can reduce strain around the implants, especially under lateral loading.

In general, an increase in implant diameter can markedly reduce bone stresses and/or strains values. Under the circumstances of immediate loading, increasing the implant diameter also decreases the bone strain around the implants. 11, 24 The IIOS454 and IIOS343 implants used in the present study have the same degree of 0.7 mm expanded platform-switching design but a different implant diameter. The IIOS454 implant, which has a larger diameter, produced significant less bone strain than the IIOS343 implant under lateral loading. These results indicate that under the same platform-switching degree of 0.7mm, the bone strain around immediately loaded implants may still also be affected by the implant diameter, with the bone strain being lower around wider implants.

However, under vertical loading, the bone strain was lowest for the NIOSM311 implant and highest for the IIOS343 implant. It therefore appears that the designs of platform switching and widening implant diameter are not always advantageous for decreasing bone strain, especially under vertical loading. Since the underlying biomechanical mechanism remains unclear, further investigations are needed of the effects of small-diameter implants with an expanded platform-switching design on the surrounding bone strain under vertical loading.

In terms of the data related to implant stability, ITV and ISQ did not differ among the three implant models. However, PTV was lower for the IIOS343 implant than for the two other groups. This might have been due to the measurements being made at the abutment level or the test peg level of the equipment—this requires further investigation.

The limitations of this study were the simplified nature of the bone geometry, material properties of bone, and muscle-attachment boundary condition. Even though the strength of a bone block is similar to that of jaw bone (according to the ASTM F-1839 certification), the strain patterns might vary with the bone geometry, other material properties (e.g., viscoelasticity), and the realism of the muscle-attachment boundary condition. Future studies should employ more sophisticated simulations of muscle attachments and material properties of bone in order to clarify the surface strain inconsistencies. Even so, the conclusions of this study can still be regarded as general principles and hence constitute useful information for clinicians. That is, under lateral loading the strain values vary significantly with the widths of the implant and abutment (platform switching) and with the implant diameter.

Conclusions

Within the limitations of the in vitro experiments performed in this study, it can be concluded that using an expanded platform-switching design with an immediately loaded small-diameter implant can reduce the strain around the implant under lateral loading. Moreover, for the same degrees of expanded platform-switching, the bone strain around an immediately loaded small-diameter implant may also be affected by the implant diameter.

Acknowledgements

This research was supported by a Taiwanese Chang Gung Research Grant (no. CMRPG8C1191) and the National Science Council of the Taiwanese government (NSC 101-2314-B-039-022-MY3).

Conflicts of Interest

None of the authors have a financial relationship with any private companies or organizations.

References

1.	Lindhe J, Meyle J. Peri‐implant diseases: Consensus Report of the Sixth European Workshop on Periodontology. Journal of Clinical Periodontology 2008;35:282-285.

2.	Lang NP, Berglundh T. Periimplant diseases: where are we now?–Consensus of the Seventh European Workshop on Periodontology. Journal of Clinical Periodontology 2011;38:178-181.

3.	Heckmann SM, Linke J, Graef F, Foitzik C, Wichmann M, Weber H-P. Stress and inflammation as a detrimental combination for peri-implant bone loss. Journal of Dental Research 2006;85:711-716.

4.	Khraisat A, Abu‐Hammad O, Al‐Kayed AM, Dar‐Odeh N. Stability of the Implant/Abutment Joint in a Single‐Tooth External‐Hexagon Implant System: Clinical and Mechanical Review. Clinical Implant Dentistry and Related Research 2004;6:222-229.

5.	Baggi L, Cappelloni I, Di Girolamo M, Maceri F, Vairo G. The influence of implant diameter and length on stress distribution of osseointegrated implants related to crestal bone geometry: a three-dimensional finite element analysis. The Journal of Prosthetic Dentistry 2008;100:422-431.

6.	Lazzara RJ, Porter SS. Platform switching: A new concept in implant dentistry for controlling post restorative crestal bone levels. The International Journal of Periodontics & Restorative Dentistry 2006;26:9-17.

7.	Vela-Nebot X, Rodríguez-Ciurana X, Rodado-Alonso C, Segalà-Torres M. Benefits of an implant platform modification technique to reduce crestal bone resorption. Implant Dentistry 2006;15:313-320.

8.	Cocchetto R, Traini T, Caddeo F, Celletti R. Evaluation of hard tissue response around wider platform-switched implants. The International Journal of Periodontics & Restorative Dentistry 2010;30:163-171.

9.	Enkling N, Jöhren P, Klimberg V, Bayer S, Mericske‐Stern R, Jepsen S. Effect of platform switching on peri‐implant bone levels: a randomized clinical trial. Clinical Oral Implants Research 2011;22:1185-1192.

10.	Calvo-Guirado JL, Ortiz-Ruiz AJ, Lopez-Mari L, Delgado-Ruiz R, Mate-Sanchez J, Bravo GL. Immediate maxillary restoration of single-tooth implants using platform switching for crestal bone preservation: a 12-month study. The International Journal of Oral & Maxillofacial Implants 2009;24:275-281.

11.	Hsu J-T, Fuh L-J, Lin D-J, Shen Y-W, Huang H-L. Bone strain and interfacial sliding analyses of platform switching and implant diameter on an immediately loaded implant: experimental and three-dimensional finite element analyses. Journal of Periodontology 2009;80:1125-1132.

12.	Andersen E, Saxegaard E, Knutsen BM, Haanæs HR. A prospective clinical study evaluating the safety and effectiveness of narrow-diameter threaded implants in the anterior region of the maxilla. The International Journal of Oral & Maxillofacial Implants 2000;16:217-224.

13.	Chong L, Khocht A, Suzuki JB, Gaughan J. Effect of implant design on initial stability of tapered implants. Journal of Oral Implantology 2009;35:130-135.

14.	Misch CE. Contemporary implant dentistry. Implant Dentistry 1999;8:90.

15.	Hahn J. Clinical uses of osteotomes. Journal of Oral Implantology 1999;25:23-29.

16.	Alsaadi G, Quirynen M, Michiels K, Jacobs R, Van Steenberghe D. A biomechanical assessment of the relation between the oral implant stability at insertion and subjective bone quality assessment. Journal of Clinical Periodontology 2007;34:359-366.

17.	Steiner M, Mitsias ME, Ludwig K, Kern M. In vitro evaluation of a mechanical testing chewing simulator. Dental Materials 2009;25:494-499.

18.	Annibali S, Bignozzi I, Cristalli MP, Graziani F, La Monaca G, Polimeni A. Peri‐implant marginal bone level: a systematic review and meta‐analysis of studies comparing platform switching versus conventionally restored implants. Journal of Clinical Periodontology 2012;39:1097-1113.

19.	Luongo R, Traini T, Guidone PC, Bianco G, Cocchetto R, Celletti R. Hard and soft tissue responses to the platform-switching technique. The International Journal of Periodontics & Restorative Dentistry 2008;28:551-557.

20.	Maeda Y, Horisaka M, Yagi K. Biomechanical rationale for a single implant‐retained mandibular overdenture: an in vitro study. Clinical Oral Implants Research 2008;19:271-275.

21.	Schrotenboer J, Tsao Y-P, Kinariwala V, Wang H-L. Effect of platform switching on implant crest bone stress: a finite element analysis. Implant Dentistry 2009;18:260-269.

22.	Chen C-LCC-S. Biomechanical effect of platform switching in implant dentistry: A three-dimensional finite element analysis: National Yang-Ming University; 2010.

23.	Calvo‐Guirado JL, Gómez‐Moreno G, López‐Marí L, Guardia J, Negri B, Martínez‐González JM. Crestal bone loss evaluation in osseotite expanded platform implants: a 5‐year study. Clinical Oral Implants Research 2011;22:1409-1414.

24.	Huang H-L, Hsu J-T, Fuh L-J, Tu M-G, Ko C-C, Shen Y-W. Bone stress and interfacial sliding analysis of implant designs on an immediately loaded maxillary implant: a non-linear finite element study. Journal of Dentistry 2008;36:409-417.

Figure legends

Fig. 1. Standard-diameter (left) and small-diameter (middle) implants with an expanded platform design were compared to the small-diameter (right) implant with a regular platform design.

Fig. 2. A smaller internal-hex abutment was attached to the expanded platform of the implant to implement the concept of platform switching.

Fig. 3. A 30-degree lingual lateral force of 190 N was applied to the top of the abutment by a universal testing machine. The microstrains in the bone near the implant were measured by a data acquisition system and analyzed by LabVIEW SignalExpress software.

Fig. 4. Medians and interquartile ranges (IQRs) of the peak values of principal microstrains on the buccal (left) and lingual (right) sides of bone in the three implant groups under vertical loading.

Fig. 5. Medians and IQRs of the peak values of principal microstrains on the buccal (left) and lingual (right) sides of bone in the three implant groups under lateral loading.